I.

체온 실시간 모니터링을 위한 초소형 무선 텔레메트리 모듈

A Miniaturized Telemetry Module for Real-time Body Temperature Monitoring

Heejoon Park Ph.D.*, Junghyun Lee MS.†, Yeonkwan Moon MS.†, Hyochan Jeon MS.*,      Kijung Ju Ph.D.*, Kyungil Yoon Ph.D.*, Namhee Park MD.‡, Yoonnyun Kim M.D., Ph.D.*

* Dept. of Medical Informatics, School of Medicine, Keimyung University, Daegu 700-712, hjpark@kmu.ac.kr

†Dept. of Electrical Engineering and Computer Science, Kyungpook National Univrsity, Daegu, 702-701

‡Dept. of Chest Surgery, School of Medicine, Keimyung University, Daegu 700-712

Abstract: In order to monitor the variation of body temperature from multi-patient in real-time in the hospital room, a miniaturized wireless telemetry module was designed and implemented. The transmitting part of telemetry system consists of a temperature sensor, a miniaturized UHF transmitter with a transmitting antenna, a microprocessor for data encoding, and a battery. The length, width, and thickness of implemented transmitter are 10 mm, 22 mm, and 9 mm respectively. The receiving part consists of a receiving antenna, a RF receiver, and a processor for data decoding, and the decoded data was transformed to RS232C format to be transferred to computer. From the experimental result in the simulation environment that resembles a hospital room, the implemented telemetry module was verified to operate well with sufficient reliability.

Keyword: clinical pathway, electronic medical record, patient group, problem-oriented, clinical support system

배 경

일반적으로 병실에 입원한 환자의 경우 간호사는 활력징후 (vital sign)라 불리는 혈압, 맥박, 호흡, 체온과 같은 생체신호를 주기적으로 측정하여 기록한다. 이는 의사의 지시에 의해 수행되며 통상 1일 4회, 특수한 상황에서는 필요에 따라 비정기적으로 시행하기도 한다. 지금까지는 관습에 의해 기본적으로 1일 4회 측정하고 있으나 이는 보건의료 자원의 효율적인 운용이라는 측면에서 볼 때 다분히 낭비적인 요소를 안고 있다.

최근 전자공학의 발달과 정보통신기술의 의료분야에의 적용이 확대됨에 따라 단순한 생체신호의 측정을 자동화하고 이를 무선 환경에서 구현하고자 하는 노력이 계속되고 있다. 특히, 환자의 행동을 구속하지 않고 필요한 생체 정보를 측정하여 무선으로 전송할 수 있는 형태의 무선 진단방법에 대한 연구가 많은 연구진들에 의해 이루어져 왔다. 본 논문에서는 병실 내 무선 환경에서 다수의 환자로부터 체온 변화를 실시간으로 모니터링하기 위한 초소형 텔레메트리 시스템을 설계 및 구현하고자 하였다.

II. 방 법

제작된 텔레메트리 시스템의 송신부는 체온 측정을 위한 온도 센서, 초소형 송신기 및 송신 안테나, 데이터 인코딩을 위한 프로세서와 배터리로 구성하였으며, 가로, 세로, 및 두께가 각각 1 cm, 2.2 cm, 및 0.9 cm의 크기로 제작되었다. 수신부는 수신 안테나, RF 수신기 및 데이터 디코딩을 위한 프로세서로 구성하였으며, 디코딩된 신호를 RS232C 포맷으로 변환하여 컴퓨터로 전송하도록 구현하였다.

III. 결 과

모의 인체를 이용하여 가상의 병실 내 무선 환경을 구성하고 실시간으로 체온의 변화를 모니터링 하는 실험을 수행한 결과, 제작된 텔레메트리 시스템이 충분한 안정성을 가지고 동작함을 확인하였다.

IV. 결 론

본 논문에서는 병실 내 무선 환경에서 다수의 환자로부터 체온 변화를 실시간으로 모니터링하기 위한 초소형 텔레메트리 시스템을 설계 및 구현하였다. 초소형 송신부와 수신부를 제작하여 무선 병실과 유사한 환경에서 실험한 결과 구현된 시스템의 안정적인 동작을 확인하였다. 본 논문에서 구현된 텔레메트리 기술은 향후 의료정보시스템과 연동한 무구속 생체신호 계측 환경을 구현하는 데 필수적인 기술로 응용될 것으로 전망된다.

참고문헌

1. Cem F. Arkilic, Ozan Akca, Akiko Taguchi, Daniel I. Sessler, and Andrea Kurz, Temperature Monitoring and Management During Neuraxial Anesthesia: An Observational Study, Anesth Analg 2000;91:662-6.

2. D. Georgiadisa, S. Schwarza, S. Cencettib, and S. Schwaba, Noninvasive Monitoring of Hypertensive Breakthrough of Cerebral Autoregulation in a Patient with Acute Ischemic Stroke, Cerebrovascular Diseases 2002;14:129-132.

3. Ron Dueck, Noninvasive Cardiac Output Monitoring, CHEST 2001;120-2;339-341.